#### WBC (Leukocyte) Manual Counting

Sample:

- EDTA anticoagulated whole venous blood.

Preparation of Diluting Fluid:

- 2% Acetic Acid  in distilled water ( Turk’s solution)  (2 ml  glacial  acetic  acid

+ 98 ml distilled water).

Procedure:

1- Draw blood by WBC pipette to mark 0.5 and diluted the blood by Turk’s  solution to the mark 11.

2-After mixing blood gently for about 3 minutes, the first 3-4 drops discarded and   a drop of diluted blood placed into the hemocytometer chamber after cleaning and  fixation of   special cover slid, then allowed to settle for 1 minute.

3-Examine the chamber under low power, counting the total number of  WBC in four large squares including 16 medium squares.

Calculations:

Total WBC Count = N/4*20*10=N*50

Where:

- N = Total WBC counted by the counting chamber.

- Large squares=4

- Depth factor in mm =  10

- DF = Dilution Factor = 20

Red blood cells count:

1- Draw blood by RBC pipette to mark 0.5 and diluted the blood by  Hayem`s solution to the mark 101.

2-After mixing blood gently for about 3 minutes, the first 3-4 drops discarded and   a drop of diluted blood placed into the hemocytometer chamber after cleaning and  fixation of   special cover slid, then allowed to settle for 1 minute.

3-Examine the chamber under low power, counting the total number of RBC in five medium squares including the four corners and a  center one.

RBCs count=(N/80) x 400x200x10=(  ) cells/ μl

= N x 10000=(  ) cells/ μl

N=number of cells in 5 medium squares (80 small squares).

N/80= average number of cells in counting smalls squares.

400=total number of squares (16x25).

200=diluting factor (0.5=101, 1=200).

10=the reciprocal of the depth (1=10).

Morphological classification of anemia:

The blood indices; MCV, MCH and MCHC were used in morphological classification of anemia.

1-Macrocytic  normochromic.

2- Macrocytic   hypochromic.

3- Normocytic  normochromic.

4- Normocytic  hypochromic.

5- Microcytic  normochromic.

6- Microcytic  hypochromic.

MCV: low= microcytic, normal = normocytic and  high= macrocytic.

MCH and MCHC: low= hypochromic and  normal= normochromic.

MCV=PCV / RBC (million) x10(fl= femtoliter )

MCH=Hb/RBC (million) x10 = pg( picogram )

MCHC=Hb / PCVx100	=(g/dl)

# Text boxes

Begin